## Summary Template
BURKE CALIBRATION LTD.
8013 110th St
Grande Prairie, Alberta T8W 6T2
Ph. (780) 897-5869
www.burkecalibration.com
TORQUE WRENCH CERTIFICATE OF CALIBRATION
Manufacturer: | JET |  |  | Serial Number: | 23034086
Model: | JTW-12250 |  |  | Calibration Date: | 2025-03-10
Max Torque: | 250.0 ft lb |  |  | Calibration Due: | 2026-03-10
Unit #: | 718912
Customer: | Big Bear Pipetffing
Phone Number: | 403 415 6032
Address:
Applied Torque | Min - Max Allowance* | Test Results
230 | 220.8 - 239.2 | 231.2 | 225.2 | 223.7 | 223.8 | 225.2
150 | 144.0 - 156.0 | 155.7 | 154.5 | 155.5 | 154.1 | 154.5
70 | 67.2 - 72.8 | 69.5 | 69.0 | 69.4 | 69.0 | 68.9
All values are displayed in ft/lbs
* The Max and Min allowance is in accordance with ISO 6789 : 2017
Burke Calibration Ltd Certifies that the test equipment used in the performance of the above test has been calibrated using standards traceable to the National Institute of Standards and Technology (NIST). The torque measuring device satisfies the requirements of ISO 6789-1 : 2017 , ISO 6789-2 : 2017 and complies with Lab requirements of ISO 17025 : 2017
Temperature: | 21C +/- 2 Deg C
Humidity: | 40%
Technician:
Standard for Calibration: Model 43220 S/N 124172 Certificate Number 231913